Model 720 Pressure Transducer Communication Protocol

Modbus RTU One Page Quick Reference

Communication Defaults

| Slave | Baud rate | Serial format | Decimal | Unit | Byte order | CRC order |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 9600 bps | 8 data bits, 1 stop bit, no parity | 1 place | MPa | Data MSB first | CRC LSB first |

Core Command Frames

| Operation | Host Request | Transducer Response | Meaning |
| --- | --- | --- | --- |
| Read pressure | 01 03 00 04 00 01 C5 CB | 01 03 02 13 88 B5 12 | 0x1388 = 5000; apply configured decimal and unit |
| Zero adjustment | 01 06 00 05 00 01 58 0B | Echoes the complete request | Write 0x0001 to register 0x0005 |
| Read parameter | 01 03 RegAdd_MSB RegAdd_LSB 00 01 CRC_LSB CRC_MSB | 01 03 02 Data_MSB Data_LSB CRC_LSB CRC_MSB | Function 03; read one register |
| Write parameter | 01 06 RegAdd_MSB RegAdd_LSB RegDat_MSB RegDat_LSB CRC_LSB CRC_MSB | Echoes the complete request | Function 06; write one register |

Register Map

| Name | Address | Function | Data Definition |
| --- | --- | --- | --- |
| Addr | 00 00 | Transducer address | 1 to 255 |
| Baud | 00 01 | Communication baud rate | 1 = 2400; 2 = 4800; 3 = 9600; 4 = 19200 |
| Unit | 00 02 | Engineering unit | 0 = m; 1 = kPa; 2 = MPa; 3 = °C; 4 = L; 5 = bar; 6 = psi; 7 = Pa |
| Dot | 00 03 | Decimal-point position | 0 to 4 |
| Pressure | 00 04 | Measured pressure | Read-only measurement value |
| Zero | 00 05 | Zero offset adjustment | Write 1 to execute |
| Parity | 00 06 | Parity mode | 0 = none; 1 = odd; 2 = even |
| Loc | 00 0A | Password validation | Password = 38 79 |

Parameter Writing Sequence

| Step | Action | Frame |
| --- | --- | --- |
| 1 | Validate password 0x3879 at register 0x000A | 01 06 00 0A 38 79 7B EA |
| 2 | Write the required value; example changes slave address to 2 | 01 06 00 00 00 02 08 0B |

Notes: Password validation remains active until power is removed. After changing communication settings, update the host before sending the next command. Verify the target register and value before writing.