522 Portable Hydraulic Pressure Pump
Operation Manual

1. Product Overview

The 522 Portable Hydraulic Pressure Pump is a manual pressure source designed for calibration and testing of pressure measuring instruments. It features a compact structure, convenient operation, stable pressure output, and is suitable for field calibration applications.

2. Technical Specifications

Pressure Range: -0.08~60 MPa

Pressure Medium: Transformer oil or clean water

Minimum Adjustable Resolution: 10 kPa

Connection Thread: M20×1.5

Dimensions: 345 × 225 × 150 mm

Weight: Approx. 4 kg

Operating Temperature: 0~50℃

Operating Humidity: <85%

3. Main Components

Liquid Filling Port

Pre-pressurizing Handwheel (Coarse Adjustment)

Pressure Release Valve

Shut-off Valve

Fine Adjustment Valve (Pressure Increase)

Pressure Instrument Connection Port

4. Operating Instructions

4.1 Preparation

Before use, open the liquid filling port and fill with clean transformer oil, hydraulic oil, or pure water. During operation, keep the filling port slightly loosened to maintain ventilation.

4.2 Pressure Calibration Procedure

1. Connect the standard pressure instrument and the instrument under test.
2. Perform air removal by operating the pre-pressurizing handwheel.
3. Close the release valve and adjust pressure gradually using the fine adjustment valve.
4. Compare readings at each calibration point and record the results.

4.3 Pressure Reduction Calibration

Slowly rotate the fine adjustment valve to reduce pressure. Compare and record the readings at each calibration point during the pressure decreasing process.

4.4 Zero Point Calibration

Open the pressure release valve and shut-off valve to perform zero point verification.

4.5 Vacuum Calibration

Adjust the pre-pressurizing handwheel and fine adjustment valve to the required position, then operate gradually to achieve the required negative pressure point.

5. Safety Precautions

Use the hydraulic pump within the rated pressure range. Do not exceed the maximum safe pressure of 80 MPa.

When transporting or carrying the pump, lock the quick connector with plugs and fully tighten adjustment handwheels.

Keep the liquid filling port open during operation.

Do not apply excessive force to valves, knobs, handles, or connectors.

Replace the pressure medium if it becomes contaminated.

Keep exposed threaded parts clean and lubricated.

6. Application

The 522 Portable Hydraulic Pressure Pump is widely used for calibration, testing, and maintenance of pressure gauges, pressure transmitters, and other pressure instruments in industrial applications.